Olympus FluoView Virtual Channels

- Select the “Dye List” option.

- Select “Virtual Channel Scan” at the top of the Dye List.

- Select “Number of Phases Used”.  Note:  number of phases and number of dyes are not the same.

- Select “All Clear” to remove any previous dye settings.

- Double click on the name of each dye that will be used.  The software will automatically assign the dyes to a Phase.

- Select “Apply”.

- From the “Virtual Channel Controller” select Phase 1.

- Using the Image Acquisition Controller, select XY Repeat and optimize the image(s).  Once optimized, select Repeat Stop.

- Repeat steps 7/8 for each Phase.

- When all Phases have been optimized, select “Start” from the Virtual Channel Controller.  Virtual Channels can be used for a single XY images as well as XYZ and XYT image series.

- If you want to end the scan early, select “Stop” from the Virtual Channel Controller.

- Once the Virtual Channel scan ends, the multi-channel image will be displayed in an overlaid format.