Spectral FluoView Start-Up Procedure

- Turn the power strip located behind the monitors to the ON position.  This will turn on the microscope power, the confocal power and the monitors.

- Start the computer.

- Turn the Mercury arc lamp power supply ON (“I”).

- Turn Argon laser power switch ON (“I”) and turn key to the ON position (“I”).

- Turn the 559nm Laser Diode power switch ON (“I”).  When the  “temp” light stops flashing green (look for solid green light), turn the key to the ON position.  Once the key is turn, the two red “laser” lights will flash until the laser is ready to use.

Spectral FluoView Shut-Down Procedure

- Turn the Mercury arc lamp power supply (#3)  OFF (“O”).

- Turn the key to the Argon laser (#4) to the OFF (“O”) position.  Wait 5 minutes for the laser to cool, and then turn the Argon laser power switch to the OFF (“O”).

- Turn the key on the 559nm Laser Diode (#5) to the OFF position.  Then, turn the power switch to the OFF (“O”) position.

- Shut-down the computer software.

- Turn the power strip (#1) located behind the monitors to the OFF position.